Předmětem této VZ je laserový litografický systém zaměřený na přípravu mikrostrukturovaných povrchů pomocí techniky přímé laserové litografie.

Minimální technické parametry zařízení:

- Pracovní rozsah laseru min. 140x140 mm²

- Možnost upevnění substrátů (waferů) až do X-Y 200x200 mm2 a tloušťky min. 300µm – 5mm

- Nainstalovaný vakuový držák vzorku pro wafery o průměru až 6"

- X-Y rozlišení stolku min. 150 nm nebo vyšší

- Systém musí být schopen pracovat s malými vzorky 3x3mm² se stejnou kvalitou jako na celém waferu.

- Instalován bude laser 375nm, s min. výkonem 60 mW (životnost >10 000h)

- Možnost integrovat více různých laserových zdrojů v jednom přístroji: např. 375nm, 405nm (nebo 266 nm). Tyto lasery musí být možné manuálně nebo automaticky přepnout během několika sekund (nabízená konfigurace musí být upgradovatelná na nejméně 2 simultánně používané lasery.

- Možnost využití více velikostí laserových spotů. Minimálně 2 velikosti, které jsou automaticky přepínatelné během procesu, musí být obsaženy v dodaném zařízení.

- Systém musí být vybaven 1ks optické trubice s velikostí laserového spotu < 1μm a 1ks optické trubice s velikostí spotu 10μm

- Systém musí být schopen automatického přepnutí optické dráhy (1μm/10μm) během litografického procesu

- Systém musí mít dostatečnou hloubku ostrosti, která dovoluje litografický zápis na silné vrstvě fotorezistu během jednoho průchodu laseru bez nutnosti zaostřování.

- Minimální aspect ratio zařízení 1x20 nebo více

- Schopnost zápisu na rezist o síle 250µm nebo silnější

- Mód zápisu: skenovací mód, vektorový mód a kombinace obou módů v jednom vzoru.

- Systém by měl být schopen poskytnout min. 5000 odstínů šedi

- Systém musí být vybaven overlay funkcí pro všechny typy vzorků s přesností 1µm nebo lepší

- Systém musí být dodán s ovládacím PC a kontrolní jednotkou se softwarem pro generování mikrovzorů (kompatibilní DXF-GDSII) a kontrolním softwarem pro laserový litograf

- Automatická fokusace

- Antivibrační pneumatický stolek a další antivibrační komponenty musí být součástí dodávky zařízení

- Systém musí být kompatibilní s čistými provozy.

The subject of the tender is the laser lithography system dedicated for preparation of microstructured surfaces by direct laser lithography method.

The minimal technical specifications of the equipment:

- Laser working surface min. 140x140 mm²

- Possibility to mount substrates (wafers) up to X-Y size 200x200 mm2 and with min. 300µm – 5mm thickness

- Installed vacuum chuck for wafer till 6"

- X-Y stage resolution min. 150 nm or higher

- System must be capable to proceed small sample 3x3mm² with the same rendering quality than wafers

- Installed laser source 375nm, min. power 60 mW (life-time >10 000h)

- Possibility to integrate several different laser sources in the machine: 375nm, and 405nm (or 266 nm) laser for example, switchable manually or automatically in few seconds (offered system must be upgradable at least up to 2 different simultaneously installed lasers).

- Possibility to use several beam sizes, minimum 2 beam sizes switchable automatically during the patterning with the same device.

- System will be equipped with 1pc of optical tube with the beam size < 1μm and 1pc of optical tube with the beam size 10μm

- System must be able to perform the automated optical line switch (1μm/10μm) during the lithography process

- System must have a large depth of focus allowing the writing into thick resist layers in one laser pass without focusing change.

- Minimum aspect ratio 1x20 or more

- Capability to write into 250µm thick layer or more

- Mode of laser writing: scanning mode, vector mode and combination of both for the same pattern

- System should be able to work with min. 5000 greyscale levels

- System must be equipped with the overlay function with better than 1µm accuracy on all kind of samples

- System must be delivered with the control PC unit and appropriate software for design of micropatterns (compatible DXF-GDSII) and control software for lithography equipment

- Automated focus settings

- Antivibration pneumatic table or other types of antivibration systems must be delivered with the equipment

- System must be clean room compatible